## Slide 1: Team 50: Remote Biometric System Bi-Weekly Update 4 Simon Becker Andrew-Joseph Nicolas Lance Pagaduan Harin Sankar Sponsor: Dr. Steven Wright TA: Hailey Ward
Notes: Michelle

## Slide 2: Integrated System Diagram Hardware
Patient MRI Data
Cardiac Data
Software
Radar
PCB
Data Capture Board

## Slide 3: Integrated System Diagram Software

## Slide 4: Project Timeline
Subsystem Designs and Testing (completed 2/2) | PCB Ordered (completed 2/4) | MCU and PCB Functionality Tests (2/26) | Radar & Data Capture Card Integrated (03/04) | MCU and PCB Deployable (to complete by 3/16) | Integration of Scan Filter and MRI (to complete by 3/16) | Integration of Scan Filter and Desktop Application (to complete by 3/16) | System Testing (to complete by 3/23) | Validation (to complete by 4/6) | Demo and Report (to complete by 4/19)
Notes: Eric

## Slide 5: PCB Andrew-Joseph Nicolas
Accomplishments since last update 25 hrs of effort | Ongoing progress/problems and plans until the next presentation
Board fully soldered and tested. All rails work as expected. MCU flashed successfully and passed blink test. | Fully integrate and validate the MCU subsystem on PCB.

## Slide 6: PCB Andrew-Joseph Nicolas
MCU current draw:
Successful communication:
Flash confirmation:
DC Rail: | Expected Voltage Rating: | Actual Voltage Rating (V)
5V @ 2A | 4.985 | 4.978
5V @ 3A | 4.991 | 4.985
3.3V @ 1A | 3.297 | 3.296
Rail Voltage Measurements:

## Slide 7: Radar/MCU Lance Pagaduan
Accomplishments since last update 13 hrs of effort | Ongoing progress/problems and plans until the next presentation
Radar outputs accurate and consistent heart rate and respiration rate with adjusted configuration file Radar and data capture card integrated | Begin live testing of raw I/Q data Test MCU code with collected I/Q data to verify correct

## Slide 8: Radar/MCU Lance Pagaduan
Radar Testing Changes
Configuration File
Testing Parameters
Patient had no loose clothing
Radar remained in a 1.5m range from the patient while elevated to chest height
Patient remains in an upright position and did not talk during data collection
Changed the frame period from 90 ms to 70 ms to increase data which updates heart rate more frequently for more accurate readings

## Slide 9: Radar/MCU Lance Pagaduan
Normal Resting Heart Rate Test

## Slide 10: Radar/MCU Lance Pagaduan
Holding Breath Test

## Slide 11: Radar/MCU Lance Pagaduan

## Slide 12: Scan Filter Harin Sankar
Accomplishments since last update 15 hrs of effort | Ongoing progress/problems and plans until the next presentation
MCU Integration Simulated whole-file cardiogram transfer and verified ingestion. Simulated live cardiogram streaming and verified ingestion. | MCU Integration Test Python cardiogram ingestion using the actual MCU hardware. MRI Integration Test k-space (.mrd) scans with python reader

## Slide 13: MRI-Scan Filter Integration Harin Sankar

## Slide 14: MRI-Scan Filter Integration Harin Sankar
Row Transfer
File Transfer

## Slide 15: User Interface Simon Becker
Accomplishments since last update 10 hrs of effort | Ongoing progress/problems and plans until the next presentation
Passed Python.NET testbench | Import Scan Filter code into the Python.NET bridge Testing the output of Scan Filter compatibility with Image Viewer

## Slide 16: User Interface Simon Becker
Python.NET Testbench:
Key Functionalities:
  Generate DICOM images with metadata
  Validate potential MCU inputs (.csv)

## Slide 17: Execution Plan
Notes: Michelle

## Slide 18: SYSTEM Validation Plan
Notes: Michelle

## Slide 19: Thank You