Radar Setup

Option 1: Radar on wall

Option 2: Radar on stand (height adjustable)

- Radar mounting options discussed:

  - Wall mount is the safer baseline option

  - Possible to use a height-adjustable nonmetal stand only if made from safe materials and approved for the MRI environment

  - Do not hand-hold the radar during scanning

  - Avoid normal metal tripods, lab stands, spring clamps, or magnetic mounts

Phantom Setup

- Main testing goal: create a small, repeatable moving phantom that fits inside the 20 cm bore and produces controlled motion so the MRI scan shows motion corruption that our project can detect and correct.

- General phantom requirements / restrictions:

  - Must fit comfortably inside the 20 cm diameter bore

  - Should be small, simple, and repeatable

  - Should use nonmetallic / nonmagnetic materials

  - Should avoid metal springs, screws, clamps, and magnetic parts

  - Should avoid air gaps / bubbles when possible

  - Should be easy to place on a sled or holder

  - Motion should be small and controlled, not large or random

  - Phantom should return to the same original position each time if possible

- Best feasible phantom option: small plastic water bottle or vial filled with water

  - Easiest and lowest-work option

  - More uniform and repeatable than fruit

  - Should be filled almost completely with water

  - Should have little to no trapped air

  - Can be mounted on a small plastic sled

- Feasible quick-test phantom option: small tangerine / clementine / orange

  - Can work for a rough early test

  - Easy to get and small enough for the bore

  - Not ideal because it is not uniform, has segments / peel / natural variation, and may contain tiny air gaps

  - Less repeatable and harder to defend as a clean engineering phantom

- Not preferred phantom option: tennis ball filled with water

  - Feasible in theory, but not a good choice

  - Requires cutting, filling, resealing, and removing air

  - Starts with a large air-filled hollow inside

  - Fuzzy outer surface makes mounting and motion less clean

  - More work and less controlled than a small water bottle

- How to create motion:

  - Place the phantom on a small plastic sled / tray

  - Sled moves along a simple plastic track

  - Motion can be along the bore axis horizontally

  - Track should guide the phantom so it moves a small amount without wobbling

- Feasible reset / return methods discussed:

  - Manual reset: simplest option; pull to one side, then manually place back before next scan

  - Gravity return: put track on a slight incline so phantom slides back when string is released

  - Double-string method: one string pulls forward, one pulls back

  - Elastic return: elastic on opposite side pulls phantom back, but may introduce bounce

- Best motion method for least work:

  - Thin nylon fishing line attached to the sled

  - Pull from the MRI opening to shift the phantom a few millimeters

  - Keep motion small and repeatable

- String / pull restrictions:

  - Use thin nonmetallic nylon string

  - Keep the string out of the imaging region as much as possible

  - Avoid letting the string shake, vibrate, or cross through the radar line of sight unnecessarily

  - The string itself is usually less of a problem than uncontrolled pulling motion

Scan Parameters

Radar-MRI Synchronization (TTL?…Other options)

- Use one shared scan-start reference so both datasets have the same “time zero.”

- Timestamp the radar cardiogram continuously during the full scan.

- Record MRI timing parameters such as TR, acquisition order, and total number of lines/frames.

- Compute MRI acquisition times from the scan parameters so each k-space line or data block has an estimated timestamp.

- Use any MRI trigger pulse if available as a cleaner alignment marker.

- Create one intentional motion event during testing that is visible in both the radar signal and the MRI data, then use it as a check.

- Keep the phantom motion repeatable so alignment errors are easier to detect.

- Use the same system clock / logging clock when possible for all saved timing data.

- Check that the radar peaks line up with the expected MRI motion-corrupted periods after alignment.

- Validate alignment numerically by measuring timing mismatch and making sure it stays within an acceptable error range.